# Figure legends

## Fig. 1. Prospective LoD protocol and two-arm replicate planning (panels A–D)

Operating characteristics for two-arm hit-rate comparison at verification concentration c*, with logistic dose–response calibration (Supplement §S2c; default γ=2 per log10 decade; Methods).

A. Three-step planning banner (lock the claim, score n=20, plan n* as a function of delta p); heatmaps use the integer (pA,pB) lattice: comparative claim θ2=pA(c*)−pB(c*); operating-characteristic surfaces πFisher(pA,pB;n) and design surface n*(pA,pB) (Eq. 2); banner links analytical LoD​95 (Pr(miss∣c*)=1−p0=0.05) to finite-n inferential OC (α, β). Concentration map c/LoD95 versus p for every integer p∈{50,…,99}% (logistic dose–response; Supplement §S2c).

B. Exact Fisher miss probability β=1−πFisher at n=20 per arm over all valid integer pairs (1225 cells; α=0.05). 61% of cells have β>0.8 (a property of the equally spaced design lattice, not a laboratory prevalence). Eight reference pairs marked (Table 2). Dashed contours: β=0.8 under α=0.01 and 0.10 (asymptotic pooled-z).

C. Design surface n*(pA,pB) for power 1−β=0.80 (Eq. 2; log color scale). Marginal by contrast, median n* versus contrast Δ=pA−pB (Supplementary Table 3f). Dashed contours: median n* at power 0.70 and 0.90.

D. Map from hit rate p to c/LoD95 under logistic dose–response (Supplement §S2c; γ=2.0 default; band γ∈[1.5,2.5]). All integer p∈{50,…,99}% shown; p=0.80 and p=0.95 annotated as examples.

Full lattice: online Supplementary Table03_comparison_grid_n20.csv; marginal by Δ: Table 3f.

Alt text (Fig. 1). Four-panel figure. Panel A: lock the claim, score n=20 on the design lattice, then plan n-star as a function of delta p. Panel B: heatmap of miss probability beta at n equals 20 for all valid p_A greater than p_B pairs; 61 percent above 0.8. Panel C: log-scale design surface n star; 90/80 pair marked. Panel D: logistic map from every integer hit rate to concentration ratio relative to LoD95.

## Fig. 2. Tier A public-deposit calibration (panels A–C)

Calibration of planning outputs on archived tier A public deposits (Table 3). Illustrates gate and guidance-n behavior; not multi-assay clinical validation of commercial assays or IFUs.

A. Mus wild-type qPCR hit-rate series near LoD (Zenodo 12609243; Piaggio et al.): standard-quantity (SQ) versus detection rate with binomial error bars; horizontal dashed line at p=0.95. Steep slope near LoD motivates locking both c* and the detection claim before planning n.

B. Ratios of guidance n to conventional n=20 for tier A deposits (deposit labels match Table 3 abbreviations). Bars plot only when a guidance n and ratio are defined; deposits without a comparable ratio are labeled “ratio N/A” (see Table 3). Where defined, values commonly fall between about 2.6× and 10× (examples: multiplex respiratory qPCR lot comparison, ≈9.95× with guidance n=199; EP17 estradiol benchmark confirmation versus qualitative pass rule, ≈2.65× with guidance n=53).

C. Equal-weight decision risk ledger comparing conventional n=20 designs with powered plans for reference protocol-specific pairs (median Lconventional/Lpowered=3.9; Table 3).

Alt text (Fig. 2). Three-panel figure. Panel A: Mus WT qPCR hit rate versus standard quantity on log scale, steep rise near LoD, dashed line at 95 percent detection. Panel B: horizontal bar chart of guidance n divided by 20 for tier A deposits where ratio is defined; some rows marked ratio not applicable. Panel C: decision-risk comparison bars for conventional n equals 20 versus powered sample sizes.