# Tables

## Table 1. Detection claim structures mapped to sample-size methods

| Goal | Statistical target | Typical protocol language | Sample-size method |
| --- | --- | --- | --- |
| Benchmark confirmation | p(c*) >= p0 | Benchmark confirmation: hit rate at c* meets IFU benchmark (typically p0=0.95); one-sided Clopper-Pearson limits | one-proportion vs p0 |
| Two-arm comparison | p_A(c) - p_B(c) | Two-arm contrast at c*: lot/reagent discrimination theta2; exact Fisher primary (Boschloo optional confirmatory) | two-proportion exact Fisher-Boschloo |
| Qualitative pass criteria | P(X>=k) | Qualitative pass rule (e.g. 19/20): operating event {X>=k}; binomial pass probability distinct from benchmark confirmation | binomial cumulative search |

## Table 2. Two-arm operating characteristics: eight reference pairs at n=20 per arm

Full integer 50–99% comparison grid (1225 valid cells): online Supplementary Table 3 CSV extension.

| (pA,pB) | Δp | n/arm | Exact Fisher power | Miss β | n* (80% power) |
| --- | --- | --- | --- | --- | --- |
| 90/80 | 0.100 | 20 | 0.072 | 0.928 | 199 |
| 95/85 | 0.100 | 20 | 0.070 | 0.930 | 141 |
| 90/70 | 0.200 | 20 | 0.242 | 0.758 | 62 |
| 90/50 | 0.400 | 20 | 0.742 | 0.258 | 20 |
| 95/90 | 0.050 | 20 | 0.017 | 0.983 | 435 |
| 80/70 | 0.100 | 20 | 0.062 | 0.938 | 294 |
| 99/95 | 0.040 | 20 | 0.002 | 0.998 | 285 |
| 85/75 | 0.100 | 20 | 0.067 | 0.933 | 250 |

## Table 3. Tier A public-deposit calibration meta-summary

One row per evidence-tier-A dataset (n=10): parametric gate outcome, model path, guidance n versus conventional n=20, ratio nguidance/n20 when applicable, and LoQ@CV20% when a precision profile applies. Blank ratio cells indicate no comparable guidance path or missing contrast inputs (Fig. 2B labels these “ratio N/A”). Tier C/S deposits are excluded by design. Supports Fig. 2.

| No. | Source | Result family | Model path | Guidance n | Conventional n | nguidance/n20 | LoQ@CV20% |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Piaggio et al. (2024) | qpcr_cq_continuous | nonparametric_percentile | 199 | 20 | 9.950 |  |
| 2 | Piaggio et al. (2024) | qpcr_cq_continuous | nonparametric_percentile | 199 | 20 | 9.950 |  |
| 3 | Analyse-it EP17 workbook | immunoassay_continuous | nonparametric_percentile | 53 | 20 | 2.650 |  |
| 4 | Analyse-it EP17 workbook | probit_hit_rate |  | 199 | 20 | 9.950 |  |
| 5 | Carobene et al. (2026) | clinical_chemistry_multilab | nonparametric_percentile | 53 | 20 |  |  |
| 6 | MSstatsLOBD / CPTAC | proteomics_blank_lob | parametric_lognormal | 53 | 20 | 2.650 |  |
| 7 | Malkamäki et al. (2024) | radioimmuno_precision_profile | parametric_normal | 53 | 20 | 2.650 |  |
| 8 | Malkamäki et al. (2024) | radioimmuno_linearity_continuous |  |  | 20 |  |  |
| 9 | plos_bzdok2024_trec_lod | hospital_qpcr_lod_replicates | parametric_normal | 199 | 20 | 9.950 |  |
| 10 | Sawant et al. (2025) | rmn_microneutralization_continuous | nonparametric_percentile | 53 | 20 | 2.650 |  |